## Slide 1: ORH128
Documentation v1
John Hermiz
09/27/17

## Slide 2: Overview
The objective of this document is to describe the features and details needed to operate the ORH128 board

## Slide 3: What is the ORH128?
It is a headstage that can record up to 128 channels of bio-potentials.
It uses a uni-polar configuration for recording signals.
  That is all the signals are recorded with respect to 1 reference
  There are several options for reference
It features four 32-channel Intan RHD2132 ASICs
Front Side
Back Side
ORH128 Board
RHD2132

## Slide 4: Inputs
Zif Connector
  Apply the bio-potentials to a ribbon cable, which mates with the zif connector
Auxiliary channels
  Two auxiliary channels are provided. These signals can be applied to the green touch proof connectors on the back side.
    Note the switches on the back side will need to be configured so that they are passed through to the amplifier
Reference and Ground
  External reference and ground can be supplied to the headstage. Use the red touch proof for external reference and black for ground

## Slide 5: Outputs
SPI Outputs
  The digital outputs from the amplifer chip (RHD2132) are sent through the 12-pin omnetics connector.
    For more details on how to communicate to RHD2132, please see Intan documentation.
    Most users do not need to worry about this.
Reference and Ground
  The selected reference can be measured by connecting the white touch proof.
  Again, ground can be accessed via the black touch proof

## Slide 6: Reference Switch
To select which reference to use, configure the switches on the front side to the desired position.
The up position means that the nodes are open and the down position means that the nodes are connected.
Pin # | Position | Description
1 | Down | Z1 to reference
2 | Down | Z2 to reference
3 | Down | Z131 to reference
4 | Down | Z132 to reference
5 | Down | External reference to reference
Note: “reference” = REF_SEL, and “external_reference” = REF_EXT
Reference Switch Block
Z1, Z2
Ref
Ref
Z131, Z132
Pin: 1 2 3 4 5

## Slide 7: Auxiliary Switch
The auxiliary switch block allows the user to connect external channels to the amplifier via the green touch proofs
Pin # | Position | Description
1 | Down | CH1 to amplifier
2 | Down | CH2 to amplifier
Pin # | Position | Description
1 | Up | Z3 to amplifier
2 | Up | Z4 to amplifier
CH1
CH2
Pin: 2 1

## Slide 8: Electrode Map
RHD2132
A
B
C
D
Ports:
REF1
REF2
REF3
REF4
L1
L2
L3
L64
L63
L62
R1
R2
R3
R64
R63
R62
L1
L64
R64
R1
REF1
REF2
REF3
REF4
See spreadsheet for detailed mapping